# Lead(Pb) and Mercury(Hg) detection Electrochemical of purification device

The PalmSens EmStat Pico module and Bismuth-film modified screen-printed electrodes (SPEs) are well-equipped for detecting lead (Pb) and mercury (Hg). This setup uses a highly sensitive and environmentally friendly electrochemical method.

Here are a detailed breakdown of the components and the detection process.

Components of the project

- PalmSens EmStat Pico Module:

A miniature, dual-channel potentiostat that controls the electrochemical measurements.

- Interface: Connects to a microcontroller (like an Arduino) or a PC via a standard UART communication protocol.

- Dual-channel: Allows you to run two independent experiments, or a single experiment with two working electrodes for advanced analysis.

- Metrohm DropSens C013 SPE:

The sensor used for detection.

  - Working Electrode (Carbon + Bi film):

The site where the electrochemical reactions (stripping and deposition) occur. The bismuth (Bi) film, an eco-friendly alternative to toxic mercury, forms an alloy with the lead and mercury, enhancing sensitivity.

  - Reference Electrode (Ag/AgCl): Provides a stable, known potential for the system to reference during measurement.

  - Counter Electrode (Carbon): Completes the electrical circuit, allowing current to flow during the measurement.

- Sample Cell: The container where the sample solution and the SPE are placed for measurement.

Detailed detection method: Anodic Stripping Voltammetry (ASV)

ASV is the most common technique for detecting heavy metals like lead and mercury, which the EmStat Pico fully supports. The process involves three main steps:

Step 1: Conditioning and sample preparation

- Prepare the electrolyte:

Create a sample solution containing the heavy metal ions (Pb2+𝑃𝑏2+and Hg2+𝐻𝑔2+)  and a supporting electrolyte, often a buffered solution like acetate buffer.

- Add bismuth:

Depending on the SPE, you may need to add a small amount of bismuth salt (e.g., Bi3+𝐵𝑖3+) to the electrolyte solution. The potentiostat will then plate the bismuth onto the working electrode along with the heavy metals. This method is called "in-situ plating" and is known to be very effective.

Step 2: Preconcentration (deposition)

- Apply a negative potential:

The EmStat Pico applies a sufficiently negative potential (e.g., -1.0 V vs Ag/AgCl) to the working electrode for a set period of time (e.g., 60–120 seconds).

- Plate the metals:

During this step, the lead and mercury ions are reduced from the solution and deposited onto the bismuth-film modified working electrode, forming a metal alloy.

  - Pb2++2e-→Pb(Bi)(Alloy with Bismuth)

  - Hg2++2e-→Hg(Bi)(Alloy with Bismuth)

Step 3: Stripping (detection)

- Apply a positive potential:

The EmStat Pico sweeps the potential in a positive direction.

- Measure the current:

As the potential increases, the deposited lead and mercury are re-oxidized and "stripped" back into the solution. This process generates sharp, distinct current peaks at characteristic potentials for each metal.

  - Pb(Bi)→Pb2++2e-

  - Hg(Bi)→Hg2++2e-

- Use SWV or DPV:

For enhanced sensitivity and a lower detection limit, use techniques like Square Wave Voltammetry (SWV) or Differential Pulse Voltammetry (DPV), which measure the difference in current at specific points in the potential sweep.

- Quantify the results:

The peak height or peak area of the resulting current is directly proportional to the concentration of the metal in the original sample. This allows for highly accurate quantitative analysis.

Why this setup is effective

- High Sensitivity:

The stripping voltammetry technique can detect heavy metals at very low concentrations, in the parts-per-billion (ppb) or even parts-per-trillion (ppt) range.

- Environmentally Friendly:

Bismuth films are a non-toxic alternative to traditional mercury electrodes, making them safer for use.

- Portability and Low Cost:

The combination of the compact EmStat Pico module with mass-produced, disposable SPEs allows for a low-cost, portable, and versatile system for on-site monitoring.

- Simultaneous Detection:

By choosing the appropriate voltammetry technique, you can detect and quantify both lead and mercury from the same sample in a single run. The metals will generate separate current peaks at different potentials.