# Manganese(Mn) and Magnesium(Mg) detection Electrochemical of purification device

System 1: Lead (Pb) and Mercury (Hg) detection

This system is an electrochemical detection setup using a specialized potentiostat and screen-printed electrodes. It is a robust and sensitive method for heavy metals.

Components:

- EmStat Pico Module: This is a dual-channel potentiostat that applies and measures electrical potential and current.

  - Interface: Connects to a main controller, like an ESP32, via UART.

- Metrohm DropSens C013 SPE: These are disposable electrochemical sensors.

  - Electrodes: Include a carbon working electrode modified with a bismuth film, a carbon counter electrode, and a silver/silver chloride (Ag/AgCl) reference electrode.

- Sample Cell: A 10mL beaker for holding the sample and the SPE.

Detection method: Anodic Stripping Voltammetry (ASV)

- Preconcentration: A negative potential is applied, causing Pb²⁺ and Hg²⁺ ions in the sample to be reduced and deposited onto the bismuth-film modified working electrode.

- Stripping: The potential is then swept positively, causing the deposited metals to re-oxidize. This creates sharp, measurable current peaks.

- Quantification: The height of the current peaks corresponds to the concentration of Pb and Hg in the sample.

- Signal processing: A Square Wave Voltammetry (SWV) technique is typically used for maximum sensitivity with the EmStat Pico.

System 2: Manganese (Mn) and Magnesium (Mg) detection

This system is a custom setup for measuring the ionic concentrations of Mn and Mg using Ion-Selective Electrodes (ISEs). This method measures voltage, not current, and requires very high input impedance.

Components:

- Manganese (Mn) ISE (Orion 9502) and Magnesium (Mg) ISE (Orion 9502BNWP): These electrodes generate a voltage change proportional to the logarithm of the ion concentration.

- Ag/AgCl Double Junction Reference Electrode: Provides a stable reference potential against which the ISEs' potential is measured.

- High-Impedance Buffer (Op-Amp based): A crucial circuit that buffers the high-impedance output of the ISEs. The LMC6001 op-amp is ideal for this with its very high input impedance (>10¹³ Ω), which prevents any current from flowing and affecting the measurement.

- ADS1115 16-bit ADC Module: The ADS1115 converts the analog voltage from the buffer circuit into a 16-bit digital signal for the microcontroller.

  - Interface: Communicates with the ESP32 via I²C.

- ESP32 DevKit V1: Acts as the central controller, reading the digital values from the ADS1115 via I²C.

- DS18B20 Temperature Sensor: ISE measurements are temperature-dependent, so this sensor provides data for temperature compensation to ensure accurate readings.

- Sample Beakers: 100mL glass beakers for each ISE.

Detection method: Potentiometry

- Sample preparation: The ISE and reference electrode are placed in a sample solution, and the resulting potential difference is measured.

- Buffering: The high-impedance buffer circuit receives the voltage from the ISE, preventing signal degradation.

- Digitization: The buffered analog voltage is passed to the ADS1115 ADC, which converts it to a 16-bit digital value.

- Microcontroller processing: The ESP32 reads the digital value via I²C. It uses a calibration curve (specific to the ISE) and the temperature data from the DS18B20 to calculate the ion concentration in ppm.