# Summary Water Monitoring & Purification Device sensors and their detection method.

This document provides an in-depth overview of the various sensors and electrochemical techniques used in modern water monitoring and purification systems. It particularly focuses on detecting pH levels and heavy metals such as lead, mercury, manganese, and magnesium using advanced sensor technologies connected through the ESP32-S3 and the EmStat Pico development kit.

1.pH Sensor

The pH sensor measures the concentration of hydrogen ions (H⁺) to determine water acidity or alkalinity. Connected to the ESP32-S3 via GPIO36, it reads analog signals and converts them into pH values ranging from 0 to 14.
It operates at 5V with an accuracy of ±0.1pH and a response time under one minute. These sensors are commonly used in water quality testing, wastewater treatment, agriculture, aquaculture, pharmaceuticals, and industrial process monitoring.

2. EmStat Pico Potentiostat Module

The EmStat Pico is a miniaturized, plug-and-play electrochemical analyzer used to detect heavy metals in water. It operates with multiple voltametric , amperometric, and potentiometric techniques such as Cyclic Voltammetry (CV) and Anodic Stripping Voltammetry (ASV). By pairing with specialized screen-printed electrodes (SPEs), it can detect trace levels of lead, mercury, and other contaminants in real-time, offering a portable and cost-effective solution for water quality analysis.

3. Heavy Metal Detection

a. Mercury Detection

- ItalSens IS-HM (Mercury-film Graphite SPE):

Detects mercury using stripping voltammetry, which deposits and then strips metal ions to produce measurable signals. It is highly sensitive, portable, and ideal for on-site water testing.

- Metrohm DropSens C013 (Bismuth-film SPE):

Uses bismuth or gold modification for mercury detection, offering a non-toxic alternative to mercury-film electrodes.

- BVT AC1.W2.R1 (Platinum-based SPE):

Requires modification with a gold electrode to detect mercury effectively since platinum alone doesn’t amalgamate with mercury.

b. Lead Detection

- ItalSens IS-HM:

Also used for lead detection via anodic stripping voltammetry (ASV). It provides high sensitivity at ppb levels but contains trace mercury, raising environmental concerns.

- Metrohm DropSens C013 (Bismuth-modified):

Offers a mercury-free alternative using bismuth film electrodes that are both sensitive and environmentally friendly.

- BVT AC1.W4.R1 (Activated Carbon/Graphite SPE):

Uses activated carbon or bismuth-modified electrodes for improved lead detection.

- OG Carbon New Silver SPE:

Detects lead after chemical modification with a bismuth film, utilizing integrated Ag/AgCl for stable reference measurement.

4. Magnesium Detection

The NT Sensors Magnesium ISE (S024) is an ion-selective electrode specifically designed to measure magnesium ion concentration in water.
It works through a selective polymer membrane that generates a voltage proportional to Mg²⁺ levels. This sensor helps assess water hardness and monitor purification efficiency, requiring calibration with standard magnesium solutions. It operates within pH 3–8.5 and 0–80°C.

5. Manganese Detection

Unlike lead and mercury, manganese detection requires more complex electrochemical methods since Mn²⁺ ions do not easily deposit on electrodes.

- BVT AC1.W4.R1 (Activated Carbon/Graphite SPE): Uses adsorptive or cathodic stripping voltammetry (AdSV/CSV) with chelating agents for manganese detection.

- MicruX Carbon S1PE: Employs manganese dioxide (MnO₂) coating or ligands such as oxine to enable manganese measurement.

- OG Carbon New Silver SPE: Uses AdCSV methods, enhanced with integrated Ag/AgCl reference electrodes for stable on-site readings.

These approaches enable trace detection of manganese, ensuring comprehensive analysis of purification performance.

6. Overall System Integration

The combination of ESP32-S3 microcontroller with the EmStat Pico potentiostat provides a versatile, IoT-based analytical system.
Each sensor — from pH to heavy metals — can be calibrated, connected, and monitored in real-time, providing valuable data for maintaining water purity standards. The document emphasizes how integrating electrochemical sensors and microcontrollers forms the foundation of intelligent, portable, and sustainable water purification systems.

Conclusion

This detailed report outlines a modern, multi-sensor water purification system that combines classic electrochemical principles with smart IoT integration. It demonstrates how modular components — like ion-selective electrodes, screen-printed electrodes, and potentiostat — can work together to ensure clean water monitoring.
Through careful calibration, sensor modification, and microcontroller interfacing, this setup provides a reliable and future-forward solution for environmental safety and public health.